General Guidelines for Maintenance to MICROCON MAP-400M

These guidelines are based upon actual usage in normal operating circumstances.

Pre-filters should be renewed about every three months, when used in a patient room. This would be based upon constant use. In areas where the unit is exposed to high traffic and great volumes of air, such as an emergency or waiting room environment, replacement may need to be performed on a monthly basis.

The HEPA filter is an extended surface high efficiency filter cell. The HEPA filter in the MICROCON was specially designed to be compatible with the impeller. Under normal conditions, you can easily expect 12 -18 months of service. Based upon the units that have been operating in the field, the general consensus is that most filters are changed on a yearly basis.

It is difficult to place a time reference for changing filters because of the many operating conditions and length of usage for each unit. Each unit is equipped with a magnehelic gauge which measures pressure differential. This is the best judge for replacement. A pre-filter should be replaced when a 1/10th of an inch increase is registered on the meter. The HEPA*should be replaced when it reaches about 75% of its initial pressure reading. Unfortunately, these provisions are more difficult to follow and quantify. Conversely, people follow the "time scenario" for the reasons outlined above.

There should be no other maintenance required on the unit other than cleaning the outside of the unit periodically.

AS A RECOMMENDED SAFETY PRECAUTION, LATEX GLOVES AND MASK CAN BE WORN WHEN RENEWING FILTERS.

*USED HEPA FILTER SHOULD BE TREATED AS CONTAMINATED SUBSTANCE.

One Industrial Way West

Building E, Unit M

Eatontown, NJ 07724

Tel: (732) 389-8922 Fax: (732) 389-8821